|  | SPECIMEN COLLECTION GUIDE Pre-Analytical Standards — Per-Test Reference (CLSI GP41 / GP44) | Doc #: SCG-F-001 Version: 1.0 Effective: ________ Review Due: ________ |
| --- | --- | --- |

SECTION A: GENERAL COLLECTION PRINCIPLES

| Principle | Standard / Requirement |
| --- | --- |
| Patient Identification | Minimum two identifiers required: full name + DOB or MRN. Confirm verbally before collection. |
| Order Verification | Confirm test order in LIS/EHR before beginning collection. Never collect without valid order. |
| Order of Draw | Blood culture → Na citrate (blue) → Serum SST (gold) → Li heparin (green) → EDTA (lavender) → Glycolytic inhibitor (gray) |
| Inversion | Invert all additive tubes immediately per manufacturer: SST 5×, EDTA 8–10×, citrate 3–4× |
| Site Selection | Avoid sites with IV lines, edema, hematoma, scarring, or lymphedema on mastectomy side |
| Labeling | Label tubes at the bedside/chair immediately after collection — before leaving the patient |
| Timing | Note exact collection date and time on requisition and tube for timed tests (cortisol, glucose tolerance, etc.) |

SECTION B: TEST-SPECIFIC COLLECTION REQUIREMENTS

| Test / Analyte | Tube Type / Color | Min Vol (mL) | Stability | Known Interferences | Special Instructions |
| --- | --- | --- | --- | --- | --- |
| CBC / Diff | EDTA (Lavender) | 3.0 | 72h @ 4°C | Clotting, EDTA-dependent agglutination | Mix 8–10×; reject clotted specimens |
| Basic Metabolic Panel (BMP) | SST (Gold/Red) | 3.5 | 8h serum @ RT | Hemolysis (K⁺, LDH), lipemia | Allow 30 min clot time; centrifuge 10 min |
| Comprehensive Metabolic Panel | SST (Gold/Red) | 5.0 | 8h serum @ RT | Hemolysis, bilirubin (ALT/AST) | Do not underfill; centrifuge promptly |
| Prothrombin Time / INR | Na Citrate (Blue) | 2.7 | 4h @ RT | Citrate ratio critical (9:1 blood:citrate) | Fill EXACTLY to line; reject under-filled |
| Lipid Panel | SST (Gold/Red) | 3.5 | Stable frozen | Lipemia (self-interfering) | 12–14h fast required; note if non-fasting |
| HbA1c | EDTA (Lavender) | 3.0 | 7 days @ 4°C | Hemoglobinopathies, recent transfusion | Note recent transfusion on requisition |
| Thyroid (TSH/T3/T4) | SST (Gold/Red) | 3.5 | 24h serum @ RT | Biotin supplementation (>5mg/day) | Instruct patient to hold biotin 2 days prior |
| Blood Culture | Blood culture bottles | 8–10 per bottle | See bottles | Contamination — strict asepsis required | Clean site 30s with chlorhexidine; allow dry |
| Urinalysis | Urine cup (plain) | 12 mL | 2h unrefrigerated | UTI bacteria overgrowth if delayed | Midstream clean catch; refrigerate if delayed |
| Urine Culture | Sterile urine cup | 5 mL | 24h refrigerated | Contamination if technique poor | Midstream clean catch; strict aseptic technique |
| Urine Protein 24h | 24h urine jug | All urine | Refrigerate during collection | Incomplete collection = invalid result | Instruct on collection procedure; provide jug |
| Serum Protein Electrophoresis | SST (Gold/Red) | 5.0 | Stable @ -20°C | Hemolysis, prolonged storage | Spin within 60 min; aliquot for storage |
| Procalcitonin | SST or Li heparin | 3.5 | 24h @ 4°C | Hemolysis, prolonged storage at RT | Process within 60 min if possible |
| Testosterone (Total) | SST (Gold/Red) | 3.5 | 7 days @ 4°C | Hemolysis, biotin | AM collection preferred (diurnal variation) |
| Cortisol | SST (Gold/Red) | 3.5 | 24h serum | Specimen timing critical | Record exact collection time; AM vs. PM matters |
| CSF Analysis | Sterile tubes (4 tubes) | 1–2 per tube | Process STAT | Cell lysis if delayed | Number tubes in order collected; process <1h |
| Body Fluid (Pleural/Peritoneal) | Sterile tubes + heparin | 3–5 | STAT processing | Clotting if not anticoagulated | Add heparin for cell count tube; culture ASAP |
| Lead (Blood Lead Level) | Special lead-free tube | 3.0 | 30 days @ 4°C | Contamination from regular tubes | Use certified lead-free collection supplies only |

SECTION C: TUBE TYPE QUICK REFERENCE

| Color | Additive | Common Use | Inversions | Notes |
| --- | --- | --- | --- | --- |
| 🔴 Red / Gold | SST Gel + Clot Activator | Chemistry, serology, immunology | 5× inversions | Allow clot 30 min; centrifuge 1500g × 10 min |
| 🟣 Lavender | EDTA | CBC, HbA1c, blood bank | 8–10× inversions | Do not freeze; mix thoroughly |
| 🔵 Blue | 3.2% Sodium Citrate | Coagulation (PT, PTT, fibrinogen) | 3–4× gentle | Fill EXACTLY to line (9:1 ratio critical) |
| 🟢 Green | Lithium Heparin | Stat chemistry, ammonia | 8–10× inversions | Process quickly; do not freeze heparinized plasma |
| ⚫ Gray | Sodium Fluoride / Potassium Oxalate | Glucose, lactate | 8–10× inversions | Fluoride preserves glucose; not for most chemistry |
| 🟡 Yellow | ACD or SPS | Blood cultures, HLA, paternity | 8–10× inversions | Keep at RT; do not refrigerate blood cultures |

SECTION D: REVIEW & APPROVAL

| Technical Supervisor | Laboratory Director | Review Date |
| --- | --- | --- |
| Signature: _______________ Name: ___________________ Date: ___________________ | Signature: _______________ Name: ___________________ Date: ___________________ | Signature: _______________ Name: ___________________ Date: ___________________ |