|  | TROUBLESHOOTING DECISION TREE AI-Driven Error Prevention & Systematic Problem Resolution Guide | Doc #: TRB-F-001 Version: 1.0 Effective: ________ Review Due: ________ |
| --- | --- | --- |

This decision tree supports Lab2Doctors' core mission: systematic elimination of pre-analytical, analytical, and post-analytical errors. Follow each branch in sequence. Document all decisions and outcomes in the relevant log (NCR, CAPA, or Incident Report).

TREE 1 — UNEXPECTED / IMPLAUSIBLE RESULT

Use when a result does not fit the clinical picture or is flagged by delta check, QC, or clinician inquiry.

| Step | Decision / Question | IF YES → | IF NO → |
| --- | --- | --- | --- |
| 1 | Was QC in range at the time of this patient result? | Proceed to Step 2 | Reject run. Repeat QC. Do not release results until QC passes. Open NCR if needed. |
| 2 | Is there evidence of specimen hemolysis, lipemia, or icterus? | Note interference on report. Repeat with fresh specimen if clinically significant. Notify clinician. | Proceed to Step 3 |
| 3 | Does the result pass the delta check criteria (DCK-F-001)? | Proceed to Step 4 | Investigate discrepancy. Verify patient ID on specimen and order. Repeat test if indicated. |
| 4 | Was the correct tube type and order of draw used? | Proceed to Step 5 | Reject specimen. Document in Specimen Rejection Log (SPC-F-001). Request re-collection. |
| 5 | Was the specimen processed within the stability timeframe? | Proceed to Step 6 | Reject or flag result. Note stability issue on report. Request re-collection if clinically required. |
| 6 | Could a medication or supplement interfere? (e.g., biotin, heparin) | Flag result with interference comment. Notify clinician. Repeat after washout if possible. | Proceed to Step 7 |
| 7 | Does repeat testing on the same specimen confirm the result? | Report confirmed result with clinical context note. Notify clinician if critical. | Escalate to Technical Supervisor. Open CAPA. Investigate instrument, reagent, and method. |

TREE 2 — QC FAILURE

Use when any QC level fails Westgard rules or falls outside acceptable limits.

| Step | Decision / Question | IF YES → | IF NO → |
| --- | --- | --- | --- |
| 1 | Did only one QC level fail (isolated failure)? | Suspect random error. Repeat that QC level once. | Multiple failures suggest systematic error. Do not run patients. Proceed to Step 3. |
| 2 | Did the repeat QC pass? | Document repeat and run patient samples. Monitor trend. | Systematic error likely. Take instrument offline. Proceed to Step 3. |
| 3 | Was a new reagent or calibrator lot recently opened? | Perform lot-to-lot comparison per Reagent Lot-to-Lot Comparison Sheet (RLL-F-001). Re-calibrate if needed. | Proceed to Step 4. |
| 4 | Has calibration verification been performed recently? | Proceed to Step 5. | Perform calibration verification (CAL-001). Recalibrate and repeat QC. |
| 5 | Has any instrument maintenance been due or missed? | Perform maintenance per PM schedule (IML). Repeat QC after PM. | Proceed to Step 6. |
| 6 | Has the QC material expired or been stored incorrectly? | Replace QC material. Prepare fresh and repeat. | Proceed to Step 7. |
| 7 | Is the problem isolated to one analyte or affects multiple? | One analyte: check specific reagent channel or probe. Multiple: suspect calibration drift or system issue. Escalate to manufacturer. | Contact Technical Supervisor. Document in Analytical Run Log. Open CAPA if recurring. |

TREE 3 — INSTRUMENT ALARM / ERROR CODE

Use when an instrument displays an error, alarm, or instrument flag.

| Step | Decision / Question | IF YES → | IF NO → |
| --- | --- | --- | --- |
| 1 | Is the error code documented in the instrument's troubleshooting manual? | Follow manufacturer troubleshooting procedure for that code. Document action taken. | Proceed to Step 2 and contact manufacturer technical support. |
| 2 | Is the alarm related to a fluid level (reagent, waste, water)? | Replenish reagent or empty waste. Prime lines. Restart run. | Proceed to Step 3. |
| 3 | Is the alarm related to a clog or probe issue? | Perform probe cleaning per maintenance SOP. Run verification sample. | Proceed to Step 4. |
| 4 | Did the error occur after a new reagent lot was loaded? | Inspect reagent for bubbles/damage. Re-prime lines. Repeat QC. | Proceed to Step 5. |
| 5 | Does the error resolve after full system shutdown and restart? | Run QC after restart. Document in maintenance log. | Escalate to Technical Supervisor. Remove instrument from service. Open NCR. Contact manufacturer. |

TREE 4 — SPECIMEN IDENTIFICATION CONCERN

Use when specimen labeling, identity, or chain of custody is questioned.

| Step | Decision / Question | IF YES → | IF NO → |
| --- | --- | --- | --- |
| 1 | Does the specimen label match the requisition / LIS order exactly? | Proceed to Step 2. | Do NOT process. Document discrepancy. Contact collection site immediately. Reject per SPC-F-001. |
| 2 | Does the result match the clinical context and prior patient history? | Report result. No further action required. | Proceed to Step 3. |
| 3 | Is the discrepancy explainable by clinical change (acute event, transfusion, medication)? | Document clinical explanation in LIS. Report with comment. Notify clinician. | Suspect specimen mix-up. Do NOT release. Escalate to supervisor immediately. |
| 4 | Can a fresh specimen be collected to confirm? | Collect and retest STAT. Hold all results pending confirmation. | Notify attending physician of concern. Document fully in NCR. Initiate CAPA. |

ESCALATION CONTACTS

| Situation | First Contact | Escalation Contact |
| --- | --- | --- |
| QC failure — cannot resolve | Technical Supervisor | Laboratory Director → Manufacturer Tech Support |
| Instrument down — patient impact | Technical Supervisor | Laboratory Director → Service Engineer (emergency line) |
| Suspected specimen mix-up | Supervisor on duty | Laboratory Director + Ordering Physician immediately |
| CAPA needed | Quality Manager | Laboratory Director for sign-off within 24h |
| OSHA-reportable incident | Safety Officer | Laboratory Director + HR (within 8h for fatality, 24h for hospitalization) |

| Developed By | Technical Supervisor | Laboratory Director |
| --- | --- | --- |
| Signature: _______________ Name: ___________________ Date: ___________________ | Signature: _______________ Name: ___________________ Date: ___________________ | Signature: _______________ Name: ___________________ Date: ___________________ |